## Slide 1: Ecom America
EN2F Emissions Analyzer

## Slide 2: ECOM America Ltd.
Established in USA in 1993
20th anniversary
Sister company of RBR-ECOM, located in Iserlohn, Germany, established in 1985
Manufacturer of high quality portable combustion / emission analyzers

## Slide 3: ECOM Products
CN
CL
EN2F
J2KN PRO

## Slide 4: Peltier cooler connected to a peristaltic pump: -Efficient gas cooling -Automatic water removal
The Advantages of Ecom
Notes: We explain to our customers the importance of proper gas cooling and ,moisture removal. And also show them how it works.

## Slide 5
Portable analyzers is all we do
Rugged packaging, robust analyzer
High strength internal pump = fast response time, consistency of results
ECOM analyzers are known for having advanced sample conditioning features
1-2 week turn-around when you send in your analyzer in for service
ECOM vs. Competition

## Slide 6: ECOM vs. Competition
Inadequate Sample Conditioning – gas coolers are generally too small to effectively cool the gas and remove the moisture.
Pump Flow Rate too low – ECOM uses a 2–3 liter per minute flow. Competitors use 0.5-1.5 liter per minute. This means:
Slow response time
Slow stabilization
Limits the use of a longer sample line
Allows gas to be in contact with moisture for longer period of time
NO2 drop-out leads to lower NOx accuracy

## Slide 7: ECOM EN2F
Up to 4 gas sensors O2, CO, NO, NO2 (True Nox)
Remote Display (optional)
Thermoelectric Gas Cooler
High Flow Pump (1.8+ lpm)
On-board Printer
Auto-Drain Water Trap
Flow meter
6-8 hour battery life
On board data storage
  – memory card slot
Bluetooth or WiFi

## Slide 8: Basic Emission Testing Procedure
1. Start the analyzer in fresh air.
2. Allow analyzer to run for 15-30 minutes to reach stable temperature.
3. Turn analyzer off and re-power. The analyzer will set accurate zero reference points & temperature compensation curves for the current temperature.
4. Insert probe into gas stream and observe the readings on the
display.
5. Upon completion of sampling, pull sample line out and allow analyzer to purge with fresh air until O2 readings are above 20.0% and other readings are below 15 ppm before powering off.

## Slide 9: Analyzer design

## Slide 10: connections

## Slide 11: KEYPAD

## Slide 12: Bluetooth remote (option)

## Slide 13: DATA COLLECTION OPTIONS
Use printer for paper print-outs (averaging or snapshots)
On-board memory card slot saves data into a .csv file
Use WiFi to connect to mobile app and save a snapshot
Use WiFi and/or USB to communicate to DAS 5.1 software
    Creates basic data reports
Advanced CRS software allows you to set-up procedures and customize data reports

## Slide 14: ANALYZER START-UP

## Slide 15: Analyzer start-up (continued)

## Slide 16: Gas analysis - choose fuel type
Once you enter Gas Analysis, the pump will turn on.
First you must select your fuel. Press up/down to scroll between the fuel type selections.
The Fuel Type selection only affects the CO2, Efficiency, and Losses calculations NOT the ppm concentration of the emissions.

## Slide 17: Gas analysis – 1-min auto zero
The 1-minute auto zero sequence does the following:
-Calibrates O2 sensor at 20.9%
-Sets zero points for all other sensors
-Sets temperature compensation curves for current temperature
-Checks mV reading of all sensors i.e. “function test”
Note: If the ambient temperature drifts more than 10°F, it is recommended to restart the analyzer to set accurate temperature compensation.

## Slide 18
After the 1-minute auto-zero sequence, you will enter Gas Analysis mode.
There are 4 display screens:
  Press up/down to toggle between 4 display screens
  1: Zoom display (4 lines – large font size)
  2: Zoom display (4 lines – large font size)
  3: Full display (9 lines)
  4: Draft/Pressure measurement (pump will turn off)
There are 4 icons at the bottom of the display:
  F1: Standby
  F2: Quick print
  F3: CO purge pump on/off
  F4: Edit display lines
Gas analysis

## Slide 19: CO purge pump
An automatic CO purge pump introduces fresh air into the CO sensor when CO concentrations exceed pre-set limits of the sensor (usually 4000ppm).
This protects the CO sensor from over-range damage.
Pressing F3 while in Gas Analysis mode activates the CO purge pump manually. The CO will dash out on the screen. Press F3 again to turn off the purge pump.

## Slide 20: Inserting the probe
MKW
Ensure that the probe tip is in the center of the exhaust gas flow
Use the temperature indicator to find the ‘core stream’
  The temperature will be the hottest in the core stream
  Watch the indicator change from up arrow to down arrow – that is the point of maximum temperature, AKA the core stream

## Slide 21: Co purge pump
An automatic CO purge pump introduces fresh air into the CO sensor when CO concentrations exceed pre-set limits of the sensor (usually 4000ppm).
This protects the CO sensor from over-range damage.
The analyzer will monitor the CO level and turn the pump off when levels drop below the maximum range of the sensors.

## Slide 22: Draft/pressure measurement

## Slide 23: PRINTing your results
Quick print (recommended): To make a print-out of the current measurements, press the F2 key located below the printer icon displayed in the lower left portion of the Gas Analysis display screen.
The printer will start and the display will indicate the status of the print job.
Print from memory: If you press the button with the printer icon, you must first press the button with the disk icon (#7) to save into temporary memory, then press the printer button. Remember to press #7 again to clear temporary memory before printing another snapshot.

## Slide 24: calculations
The calculations will only display and print out if they are valid.

## Slide 25: Adjustments
Adjustments allows the user to make changes to calculated and measured values, such as:
Units (ppm, mg/m3)
O2 Correction
Time / Date Set
Bluetooth/USB baud rates
Change info on bottom of print-out

## Slide 26: Units of Measure
Units of measure can be
selected for :
PPM
PPM – O2 Corrected
Mg/m3
Mg/m3 – O2 corrected
Mg/kWh – O2 corrected
Mg/MJ – O2 corrected

## Slide 27: General Care & Maintenance

## Slide 28: filters
The EN2F has 2 filters which should be periodically changed:
  Particulate filter (top of water trap)
  In-line smoke filter (tubing leading out of water trap)

## Slide 29: Accessories

## Slide 30: filters
The EN2F has 2 filters which should be periodically changed:
Particulate filter
  First line of defense
  Filters out particulates and NH3
  Change when it turns dark gray
In-line smoke filter
  Second line of defense
  Filters out microparticulates & smoke
  Change when it turns dark gray

## Slide 31: Batteries
The Li-Ion battery will last 6-8 hours when fully charged.
Battery condition of the analyzer and remote can be checked in the Control mode.
5.7V to 6.7V is good
Battery alarm will start after it drops below approx. 5.7V.
(If you are using the Bluetooth remote, the display will automatically go dark after a period of time to save the battery life.)

## Slide 32: Storing the analyzer
Keep the battery fully charged and operate the analyzer on AC main voltage when possible.
Fully charge the battery prior to storage or leave connected to main power if possible. Full charge takes 8-10 hours.
Store the analyzer in a cool, dry place while plugged into AC power. If it is not possible to keep the analyzer plugged in while stored, make sure to charge it overnight at least once/month.

## Slide 33: Changing printer paper

## Slide 34: Probe and sample line cleaning

## Slide 35: Calibration Mode
To enter calibration mode, select CONTROL on main menu, then OK .
Swipe calibration magnet over the area up-left of the ESC button.
After swiping the cal magnet, the analyzer will enter calibration mode.
Press up/down to position the cursor beside the sensor you are calibrating.

## Slide 36: Basic Emission Testing Procedure
1. Start the analyzer in fresh air.
2. Allow analyzer to run for 15-30 minutes to reach stable temperature.
3. Turn analyzer off and re-power. The analyzer will set accurate zero reference points & temperature compensation curves for the current temperature.
4. Insert probe into gas stream and observe the readings on the
display.
5. Upon completion of sampling, pull sample line out and allow analyzer to purge with fresh air until O2 readings are above 20.0% and other readings are below 15 ppm before powering off.

## Slide 37: Thank you
1628 Oakbrook Drive
Gainesville, GA 30507
Toll Free: 877.326.6411
Telephone: 770.532.3280
Fax: 770.532.3620
www.ecomusa.com